## Slide 1
Stockholm, 5 August 2013
Diego Real
IFIC (CSIC – Universidad de Valencia)
Proposal of a new generation of Laser Beacon for Time Calibration in the KM3NeT Neutrino Telescope
VLVnT13 - Very Large Volume Neutrino Telescope Workshop 2013

## Slide 2
2
~60 m
100 m
14.5 m
Link cables
Junction box
2
KM3NeT Collaboration
VLVnT13 Stockholm , 05-08-2013
IFIC, Valencia
UPV, Valencia
UPC, Barcelona
40 Institutes

## Slide 3
12,800 DOMs in the deep sea at 3-5 km depth
with point-to-point connection to shore
3
~60 m
100 m
14.5 m
Link cables
2500 m depth
Junction box
3
KM3NeT Neutrino Telescope
VLVnT13 Stockholm , 05-08-2013
100m
~ 860m
640 strings
18 DOM/string
11520 DOMs
Volume: ~5 km3

## Slide 4
4
~60 m
100 m
14.5 m
Link cables
2500 m depth
Junction box
4
Deployment Of KM3NeT Detection Units
VLVnT13 Stockholm , 05-08-2013
640 string with optical sensors
in the deep sea at 3-5 km depth

## Slide 5
5
5
KM3NeT Time Calibration System (I)
VLVnT13 Stockholm , 05-08-2013
The experience from the previous projects shows that a system of external light sources with a known emission time ensures the time calibration and provides measurements of the optical water properties.
Decoupling the intra/inter detection unit (DU) calibration seems the best solution.
The calibration constants are obtained putting all the information together.

## Slide 6
6
~60 m
100 m
14.5 m
Link cables
2500 m depth
Junction box
6
KM3NeT Time Calibration System (II)
VLVnT13 Stockholm , 05-08-2013
Decoupling within the same Detection Unit (intra- D.U.) and between D.U. (inter-D.U.) calibration systems
INTRA D.U. Calibration (See David’s talk):
  One Nanobeacon per Digital Optical Module. Looking upwards

## Slide 7
7
~60 m
100 m
14.5 m
Link cables
2500 m depth
Junction box
7
KM3NeT Time Calibration System (III)
VLVnT13 Stockholm , 05-08-2013
Decoupling within the same D.U. (intra- D.U.) and between D.U. (inter-D.U.) calibration systems
INTER D.U. Calibration:
  One Laser Beacon every 8 D.U
  Laser Beacons @ 532 nm
  Higher in intensity and shorter pulses < 1 ns
  More expensive but less redundancy required
  Tunable by Liquid Crystal Optical attenuator
  Collimated beam -> Diffusion device needed

## Slide 8
8
~60 m
100 m
14.5 m
Link cables
2500 m depth
8
Laser Beacon Description
VLVnT13 Stockholm , 05-08-2013
LASER BEACON:
  Titanium Container
  Internal Mechanics
  Connector
  Laser Head
  Built-in Photodiode
  Antibiofouling system
  The Voltage Controller Optical Attenuator

## Slide 9
9
~60 m
100 m
14.5 m
Link cables
2500 m depth
Junction box
9
Laser Beacon Mechanics
VLVnT13 Stockholm , 05-08-2013
Cylinder of 650 mm high and 142 mm of diameter
Made in Titanium grade 5 (Ti6AI4V) to avoid marine corrosion
Composed of three different pieces:
  The cylindrical tube
  The lower cap (connector cap)
  The upper cap with the anti-biofouling system
Designed to resist up to 5000 meters deep
Pressure test at 400 bars
Chart: Pressure
Pressure
(bars)
Time (hh:mm:ss)

## Slide 10
10
~60 m
100 m
14.5 m
Link cables
2500 m depth
Junction box
10
Internal Mechanics And Connector
VLVnT13 Stockholm , 05-08-2013
Inner mechanics:
Provides the structure where the different devices of the laser beacon are integrated
Internal mechanics developed in aluminum
two aluminum rings joined together by two aluminum beams
Connector :
Made in titanium grade 5 from MacArtney
Allows connection between the LB and the D.U

## Slide 11
11
~60 m
100 m
14.5 m
Link cables
2500 m depth
Junction box
11
VLVnT13 Stockholm , 05-08-2013
The Anti-Biofouling System
The cylinder quartz
pressure-resistant quartz
refractive index 1.4585
60 mm x 80 mm
Inserted 10 mm into the upper titanium cap
Possible to illuminate photo-sensors located at a horizontal distance of 200 m and 50 m above the seabed.
Diffuser
model 48010 from ORIEL (Flashed opal diffuser)
Thickness of 2.2 mm
Diameter of 25 mm.
Lambertian distribution

## Slide 12
12
~60 m
100 m
14.5 m
Link cables
2500 m depth
Junction box
12
VLVnT13
KM3NeT Laser Beacon Tests Head (I)
Stockholm , 05-08-2013
Diode pumped Q-switched Nd-YAG laser
Short pulses. FWHM ~ 400 ps
Energy per pulse of 3.5 µJ (manufacturer)
4.15v – 4.25 µJ measured in laboratory

## Slide 13
13
~60 m
100 m
14.5 m
Link cables
2500 m depth
Junction box
13
VLVnT13
KM3NeT Laser Beacon Tests Head (II)
Stockholm , 05-08-2013
Green @ 532 nm
Measurements carried out with the Spectrometer Ocean Optics HR4000
The laser presents the main spectrum at 532 nm as expected

## Slide 14
14
~60 m
100 m
14.5 m
Link cables
2500 m depth
Junction box
14
VLVnT13
Built-In Photodiode
Stockholm , 05-08-2013
A jitter the laser pulse emission time of a few hundred nanoseconds
Internal Built-in photodiode that provides the exact time of the laser light emission (0.5 ns)
The read-out system to acquire the photodiode signal of the laser is not located in the laser container but in the D.U where the LB is connected
The jitter of the internal photodiode has been measured using an external photodiode (Newport 812-20 photodiode with a rise time lower than 200 ps) directly illuminated by the laser beam
The FWHM of the measured built-in photodiode jitter is lower than 550 ps

## Slide 15
15
~60 m
100 m
14.5 m
Link cables
2500 m depth
Junction box
15
VLVnT13
The Voltage Controller Optical Attenuator
Stockholm , 05-08-2013
Unlike LED sources the amount of light emitted by the laser is fixed
A voltage-controlled optical attenuator using a liquid crystal variable retarder located in the beam path is used
a beam-splitting polarizing cube is used since they have a higher damage threshold to laser exposure than standard linear polarizer
The model used is optimized to work in the 420 to 680 nm range

## Slide 16
16
~60 m
100 m
14.5 m
Link cables
2500 m depth
Junction box
16
VLVnT13
KM3NeT Laser Beacon Tests
    1 laser beacon integrated in the ANTARES IL13
    1 laser beacon integrated in the KM3NeT “Nemo Tower Phase II”
Stockholm , 05-08-2013
Both use a 3.5 µJ the laser head STG-03E-1S0 from Teemphotonics which emits light with a wavelength of 532 nm after frequency doubling of the original Nd-YAG wavelength of 1064 nm

## Slide 17
17
~60 m
100 m
14.5 m
Link cables
2500 m depth
Junction box
17
VLVnT13
KM3NeT Laser Beacon Developments
Stockholm , 05-08-2013

## Slide 18
18
~60 m
100 m
14.5 m
Link cables
2500 m depth
Junction box
18
VLVnT13
KM3NeT Laser Beacon Developments
Stockholm , 05-08-2013

## Slide 19
19
~60 m
100 m
14.5 m
Link cables
2500 m depth
Junction box
19
VLVnT13
KM3NeT Laser Beacon Developments
Stockholm , 05-08-2013
New laser head. More powerful
25 µJ per pulse versus 3.5 µJ of previous head.

## Slide 20
20
20
VLVnT13
Conclusions
Stockholm , 05-08-2013
1 laser beacon integrated in the ANTARES IL13 (3.5 µJ)
1 laser beacon integrated in “ KM3NeT Nemo Tower Phase II” (3.5 µJ)
Proposal to use a new laser head for KM3NeT.
    More powerful with 25 µJ per pulse